Techinal specification For Heamodialysis Machines

| Index | Bidders specification &comments | requirements | Feartures | Subject |
| --- | --- | --- | --- | --- |
|  |  | With Troily | Dimensions (HXWXD) | Treatment Modulity & Technical data |
|  |  |  | Weight |  |
|  |  | AC 110 – 230 V | Electrical data |  |
|  |  | Watt |  |  |
|  |  | Temperature | Watter SupplY |  |
|  |  | Pressure |  |  |
|  |  | Standard AAMI |  |  |
|  |  | Sequential Dialysis | Function |  |
|  |  | Single Needle(Single Pump) |  |  |
|  |  | Double Needle Dialysis |  |  |
|  |  | TOUCh Screen display or OCM System |  |  |
|  |  | Self Cheek |  |  |
|  |  | Battery Back UP |  |  |
|  |  | UF Rate 0 to 3000 Ml/h | UF Control |  |
|  |  | Accuracy Max + 3% |  |  |
|  |  | Volometric Control |  |  |
|  |  | Chemical | Disinfectant system |  |
|  |  | Thermal |  |  |
|  |  | Combined |  |  |

Techinal specification For Heamodialysis Machines

| Index | Bidders specification &comments | requirements | Feartures |  | Subject |
| --- | --- | --- | --- | --- | --- |
|  |  | 0.1 To 10 ML/hr | Delivrey Range |  | Heparin Pump |
|  |  | 10 , 20 , 30 / ml | Syringe Size |  |  |
|  |  | Adjustable from 300to 800 ml/min With 5 Steps | Dialysate Flow meter |  | Sensors |
|  |  | 100 micro Liter/min | Bubble Detector |  |  |
|  |  | 200 to + 400- | Venous pressure |  |  |
|  |  | -300 to +300 | Arterials pressure |  |  |
|  |  |  | Blood Leackage detector |  |  |
|  |  | Adaptable to use Rutine bicarbonate Car trige in iran ( max price < 2 Euro / bag) |  |  | Should be equipped with |
|  |  | Adaptable to connecting tubes with two segment pump and blood flow monitoring |  |  |  |
|  |  | Thermometer dialysis solution (35-39)c |  |  |  |

Techinal specification For Heamodialysis Machines

| Index | Bidders specification &comments | requirements | Feartures | Subject |
| --- | --- | --- | --- | --- |
|  |  | Bicarbonate Powder | Dialysis Using | System & profiling |
|  |  | Bicarbonate Solution |  |  |
|  |  | Acetate Solution |  |  |
|  |  | ‍ sodium or Conductivity | Profiling for |  |
|  |  | Bicarbonate |  |  |
|  |  | UF |  |  |
|  |  | 50 ml/min - 500 ml/min | Blood Flow meter Adjusting system |  |
|  |  | Auto Priming system | Qualification | Standard certificate & Guardedly |
|  |  | Automatic and digital troubling detector system |  |  |
|  |  | CE or FDA | Certificate |  |
|  |  | 93/42/EEC | Standard |  |

جدول مقايسه اي ماشين همودياليز

| رديف | نام ماشين | مدل | مشخصات فني | پودر بيكربنات | قيمت (Uro) | لوله رابط | قيمت | توضيحات |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | B-Braun | ِDialog- Plus | √ | عمومي | (uro) 1058 | عمومي | 25000( ريال ) | همراه با باطري ذخيره |
| 2 | Gambro | AK96-Bio | √ | عمومي | (uro) 1058 | عمومي | 25000( ريال ) | همراه با باطري ذخيره |
| 3 | Nipro | Surdial 55 Plus | √ | عمومي | (uro) 1058 | عمومي | 25000( ريال ) | همراه با باطري ذخيره |
| 4 | Fresenius | 4008 S Calssic | √ | اختصاصي | (uro) 3/2 | اختصاصي | 30000( ريال ) | قيمت اعلام شده تقريبي مي باشد |
| 5 | J.M.S | SDS - 50 | توضيحات | اختصاصي | (uro) 16/3 | عمومي | 25000( ريال ) | فاقد شستشوي حرارتي و استاندارد CE ( براساس اطلاعات كاتالوگ 2010 ) |